## Slide 1: Hp(10) irradiations for personal dosimetry traceability
Duncan Butler
Workshop on Building Asia-Pacific Individual Monitoring Service (IMS) Capabilities toward Regional Sustainable Network.
At Pullman Bangkok King Power Hotel, Bangkok Thailand 18-20 March 2019

## Slide 2: Quantity Hp(10)
Personal dose equivalent, Hp(d)
= Absorbed dose to soft tissue at depth of d mm at a specified point in the body
Note: For weakly penetrating radiation d = 0.07 mm is of interest (“skin dose”)
For strongly penetrating radiation, d = 10 mm is of interest

## Slide 3: In calibration laboratory…
Radiation source

## Slide 4: Comment
Regulations mostly refer to “Effective dose”
This quantity includes:
assumption: radiation incident from everywhere
dose throughout the human body
organ risk factors
Very to hard to measure!
Personal dose equivalent, such as Hp(10), is an operational quantity – it can be
measured
Used as a surrogate for effective dose

## Slide 5: Irradiation to the 137Cs Beam
Direct teletherapy Cs beam (no attenuators added)
Distance 4 m
Beam diameter 40 cm
PMMA slab phantom 30 x 30 x 15 cm
2 mm PMMA added over the front of the monitors
Nominal Hp(10) rate of 35 µSv/s
Conversion from air kerma:
  1.21 Sv/Gy at 0°
  1.22 Sv/Gy at 20°
    ISO 4037 Part 3 Table 33
Nominal relative uncertainty (k=2): 6%
Traceability: Farmer Chamber calibrated at Co-60 and X-rays against ARPANSA primary standards, and interpolated to Cs-137

## Slide 6: Irradiation to the X-ray Beam
Constant potential, tungsten target X-ray tube
Beam quality 70 kVp, 15 mA
Additional filtration: 4 mm Al, HVL = 3.19 mm Al (approx. 0.11 mm Cu)
Effective energy ~34 keV for Cu
Similar to PTB therapy beam, code TH 70
Distance 3 m
Beam diameter 45 cm
Water slab phantom 30 x 30 x 15 cm of water
  outer 32 x 32 x 16.3 cm including PMMA walls
No PMMA added over the front of the monitors
Conversion from air kerma:
  1.47 Sv/Gy at 0°
  1.44 Sv/Gy at 20°
    ISO 4037 Part 3 Table 27
Nominal relative uncertainty (k=2): 6%
Traceability: Via large area transmission chamber calibrated directly against the primary standard of medium-energy X-ray air kerma

## Slide 7: What does “traceability” mean?
A primary standard of radiation dose
The radiation measurement you make
Traceability = Documented chain of measurements

## Slide 8: ISO 4037 – irradiation of personal monitors to Hp(10)
Issued in 1999
4 parts
Updated in 2019
Tells laboratories how to irradiate monitors
to Hp(10) etc

## Slide 9: ISO 4037 (1998) – irradiation of personal monitors to Hp(10)
Determine air kerma K (free in air) at reference distance (say 3 m)
Put the monitor on the ICRU slab phantom
Adjust reference point of dosimeter to be at reference distance
Cover with PMMA (2 mm for 137Cs or 4 mm for 60Co)
Use ISO 4037 conversion coefficient to obtain Hp(10) from K

## Slide 10: In the IAEA-ARPANSA comparison…
29 dosimetry services
2 beam qualities
2 angles
5 monitors per beam and angle
29 x 2 x 2 x 5 = 580 monitors to irradiate!
Have to irradiate in groups of 2 or 3
Exposure times ~ a few minutes each
Laboratory time taken = 3 days

## Slide 11: (Added uncertainty due to monitor proximity to each other)
20°
0°

## Slide 12: Uncertainty budget for primary standard laboratory for single monitor irradiation

## Slide 13: Uncertainty budget
Occupational dose
to wearer, D
Subtract background!
Primary standard
Irradiated reference monitors
Wearer
(~3 months)
Calibrate reader
Read monitor, Dmon
(What energy? What dose? How many?)

## Slide 14
Dose given to reference monitors
Variation in reading out the reference monitors
Primary or secondary standard
Uncertainty in calibration of reader
Source of uncertainty

## Slide 15
Uncertainty in dose readout
Reader calibration
Source of uncertainty
Energy response
Linearity
Fading
Statistical variation for one monitor

## Slide 16
Uncertainty in occupational dose
Dose readout
Source of uncertainty
Background
Over 3 months, can have a large variation!
Dmon - B

## Slide 17: Statistical uncertainty
5 reference monitors
All irradiated to same dose (1 mSv) by standards laboratory
Readout gives 1.2 mSv, 1.4 mSv, 0.6 mSv, 1.1 mSv, 1.2 mSv
Average = 1.1 mSv
Standard deviation = 0.3 mSv
Standard deviation of the mean = 0.3 / sqrt(5) = 0.13 mSv

## Slide 18: Statistical uncertainty
Standard deviation = 0.3 mSv
Is the estimate of the uncertainty in one reading
Standard deviation of the mean = 0.3 / sqrt(5) = 0.13 mSv
Is the estimate of the uncertainty in the average
of the 5 readings

## Slide 19
CONTACT ARPANSA
Email: info@arpansa.gov.au
Website: www.arpansa.gov.au
Telephone: +61 3 9433 2211
Freecall 1800 022 333
General Fax: +61 3 9432 1835
THANK YOU