## Slide 1: Dosimetry Standards at NMIJ/AIST Ionizing Radiation Standards Group
Researcher: T. Kurosawa, M. Kato, Y. Morishita, T. Tanaka, M. Shimizu, H.Yamaguti
J. Ishii
Technical staff for calibration service
: R. Noda, N. Matsuura
Notes: I talk about radiation dosimetry standards at NMIJ.
Our group have 5 sources, gamma ray from Cobalt and Cs, medium and low energy X-ray, beta-particle source, synchrotron radiation.

## Slide 2
Organization
NMIJ
(National Metrology Institute of Japan)
Research Institute for
Engineering Measurement
AIST(National Institute of Advanced Industrial Science and Technology)
Research Institute for
Physical Measurement
Research Institute for
Material and Chemical Measurement
Research Institute for
Measurement and Analytical Instrumentation
Radioactivity and Neutron Standards Group
Ionizing Radiation Standards Group
Group leader Tadahiro Kurosawa
Group leader Akira Yunoki
7 researchers
2 technical staffs
6 researchers
1 technical staff

## Slide 3: Dosimetry standards
Protection
  Air kerma for Cs-137, Co-60 (T.Kurosawa)
  Air kerma for low dose rate of Cs-137 (0.1~3 mSv/h) (T.Kurosawa)
  Air kerma for medium energy X-rays, 30 kV - 300 kV (T.Tanaka)
  Air kerma for low energy X-rays, 10 kV-50 kV (T.Tanaka)
  Dose equivalent of X-ray & g-ray (T.Kurosawa, T.Tanaka)
  Absorbed dose to tissue, b-rays (Sr-90/Y-90, Kr-85, Pm-147) (M.Kato)
Developing
  Dose equivalent in 3 mm depth (Kato & Kurosawa)
  Air kerma in high energy photons using a linac (T.Kurosawa)
  Diagnostics (T.Tanaka)
  Mammography (Mo/Mo, Mo/Rh, Rh/Rh, W/Rh, W/Ag, W/Al)

## Slide 4: Dosimetry standards
Therapy
  Co-60 (absorbed dose to water) (Y.Morishita)
  b-ray brachytherapy Ru-106 (M.Kato)
  High-energy photons (M.Shimizu)
  HDR brachytherapy Ir-192 (T.Kurosawa)
  Developing
  High-energy electrons (T.Tanaka)
  Proton and Carbon beams (T.Tanaka & M.Shimizu)
  Water calorimeter (M.Shimizu)
  High dose
  Developing
  Alanine dosimetry (H. Yamaguti)

## Slide 5: Facilities
X-rays
  Low energy X-rays (10 to 50 kV)
    W, Mo, Rh target
    Small free-air chamber
  Medium energy X-rays (50 to 300 kV)
    W target
    Large free-air chamber
Gamma-rays
  High dose rate room
    148 TBq Co-60, 34 TBq Cs
  Medium and low dose rate room
    185, 18.5, 3.7GBq Co-60, 222, 18.5, 1.85 GBq Cs-137
  Very low irradiation facility
    10 MBq Cs-137

## Slide 6
  Cylindrical Graphite cavity chambers
    Two 60 cc chambers, and two 6 cc chambers
Clinical beam
  Clinical Linac (Electa)
    Graphite and Water calorimeter
Beta-rays
  Beta Secondary Standard 2(BSS2)
    Pm-147, Kr-85, Sr-90/Y-90, Ru-106/Rh-106
    Extrapolation chamber
The temperature and humid are controlled.
22±1 ℃ 50 %±20 %(humidity)

## Slide 7: Equipments for standards
Electrometers for free-air, cavity and extrapolation chambers
  Vibrating-reed electrometer
  Capacitors (10nF, 1000pF, 500pF, 100pF)
Wheatstone bridge circuit for calorimeters
Thermometers, barometers and hydrometers

## Slide 8: Medium energy x-rays
Calibration Item
Free air chamber
(primary standard)

## Slide 9: High energy photons from Linac

## Slide 10: Key comparison BIPM.RI(I)-K6
BIPM graphite calorimeter
Energy | TPR20,10 | R (NMIJ/BIPM)
6 MV | 0.679 | 0.9966(47)
10, 15 MV | 0.729 | 0.9959(55)
Comparison results
NRC
PTB
LNE-LNHB
NIST
ARPANSA
NPL
NMIJ

## Slide 11: Water calorimeter
Glass cell

## Slide 12: Reference air-kerma rate for Ir-192 HDR
Cavity
chamber
Source
position
1 m
Floor
ceiling

## Slide 13: QA / QC
For QC
  Periodical check
    Gamma-rays and beta-rays
      Periodical dose rate measurements using primary chambers
      Compare the previous dose rate -> less than 0.1 %
  X-rays
    Calibrations of our woking standards(cavity chambers)
    Compare a calibration factor -> less than 0.2 %
These check include not only the check of primary standards but also other devices( electrometers, thermometers, barometers,,,,)

## Slide 14
For QA
  Join the international comparison
    BIPM.RI(I)-K6, high energy photons (2015 April)
    BIPM.RI(I)-K8, Ir-192 brachytherapy (2015 April)
    BIPM.RI(I)-K3, medium energy X-rays (2015 Sep.)
    APMP.RI(I)-K3.2013, medium energy X-rays (2016)
    APMP.RI(I)-K8, Ir-192 brachytherapy (2017)、pilot
    High energy photons (KRISS, 2016 July)
    Personal dose equivalent (ARPANSA, 2016 Aug.)
  Review by the accreditation organization and foreign experts

## Slide 15: Key Comparison (BIPM.RI(I).K3)
Metrologia 53, 06002 (2016)
Key Comparison (BIPM.RI(I)-K3) in Sep. 2015.
Key Comparison (APMP.RI(I)-K3. 2013) in Feb. 2016

## Slide 16: Number of calibration service
 | Chambers
Air kerma for gamma-rays | 17
Absorbed dose to water for Co-60 | 6
Medium energy X-rays | 8
Low energy X-rays | 4
Beta-rays | 1
High energy X-rays | 0
・NMIJ mainly calibrate secondary standards.
・There are 7 secondary laboratories for dosimeters calibration.
・We also do some irradiation service for the test of dosimeters and detectors. (samples is more than 500?)

## Slide 17: Thank you for your attention!