## Slide 1
Intercomparison on Personal Dose Equivalent (Hp(10))
for Photon fields for Individual Monitoring Service Laboratory in Southeast and East Asia Region
Vithit Pungkun

## Slide 2: Participants
9 Countries
22 Laboratories
22 Sets of dosimeter
27 Readers
27 Sets of results submitted

## Slide 3: Timeframes
Timeframes | Activities
July 2018 | Official announcement and confirmation of participation
August 2018 | participants send the dosimeters to OAP
December 2018 | Irradiation
January 2019 | Return of irradiated dosimeters
February 2019 | Participants report the result
March 2019 | OAP evaluate the Participants’ reports
March 2019 | Evaluation Workshop & Finalization of Exercise Report

## Slide 4: Dosimeters

## Slide 5: Assessment Methodology
IAEA General Safety Guide on Occupational Radiation Protection (GSG-7):
HM is the net dose measured by the participant
H1 is the exposured dose
H0 is the lowest dose that needs to be measured
  H0 is taken to be 0.08 mSv
  Value for monthly measurement regime
  Most restrictive limits

## Slide 6: OAP Secondary Standard Dosimetry Laboratory
Secondary standards for the dosimetry of
ionization radiation for Thailand
Calibration and Measurement Capabilities
(CMCs) and comparison results are listed in
the APMP Key Comparison Database
Accredited against ISO Standard 17025 to
provide exposures of personal dosemeters
traceable to the NMIJ standard

## Slide 7: Irradiation to the Cs-137 Beam
Direct Protection level Cs beam
• Distance 41m
• Beam diameter 30 cm
• slab phantom 30 x 30 x 15 cm
• Nominal Hp(10) rate of 35 μSv/s
• Conversion from air kerma:
– 1.21 Sv/Gy at 0°
Ref.Safety Reports Series No.16 page 61
• Hp(10) delivered:
– 3.00 mSv at high dose
– 0.40 mSv at low dose
• relative uncertainty (k=2): 5%
• Traceability:
A6 Ionization Chamber calibrated at Cs-137 from NMIJ primary standard
Cs-137 source
15 cm
30cm
100cm
Personal dosimeters

## Slide 8: Irradiation to the Cs-137 Beam

## Slide 9: Measurement of Exposure to the Cs-137 High dose
Exposure: 3.00 mSv
Lower Limit: 1.90 mSv
Upper Limit: 4.61 mSv

## Slide 10: Measurement of Exposure to the Cs-137 Low dose
Exposure: 0.40 mSv
Lower Limit: 0.18 mSv
Upper Limit: 0.69 mSv

## Slide 11: Irradiation to the X-ray Beam
Constant potential, tungsten target X-ray tube
Beam quality 120 kV, 10 mA
Water slab phantom 30 x 30 x 15 cm of water
Beam quality N-120 kV, 10 mA
2.00 mSv at high dose
0.40 mSv at Low dose
Beam quality N-80 kV, 10 mA
4.00 mSv at high dose
0.40 mSv at Low dose
Beam quality N-60 kV, 10 mA
2.00 mSv at high dose
0.40 mSv at Low dose
Traceability: IAEA
Cs-137 source
15 cm
30cm
100cm
Personal dosimeters

## Slide 12: Measurement of Exposure to the X-ray beam N-120 at high dose
Exposure: 2.00 mSv
Lower Limit: 1.23 mSv
Upper Limit: 3.11 mSv

## Slide 13: Measurement of Exposure to the X-ray beam N-120 at low dose
Exposure: 0.40 mSv
Lower Limit: 0.18 mSv
Upper Limit: 0.69 mSv

## Slide 14: Measurement of Exposure to the X-ray beam N-80 at high dose
Exposure: 4.00 mSv
Lower Limit: 2.56 mSv
Upper Limit: 6.12 mSv

## Slide 15: Measurement of Exposure to the X-ray beam N-80 at low dose
Exposure: 0.40 mSv
Lower Limit: 0.18 mSv
Upper Limit: 0.69 mSv

## Slide 16: Measurement of Exposure to the X-ray beam N-60 at high dose
Exposure: 2.00 mSv
Lower Limit: 1.23 mSv
Upper Limit: 3.11 mSv

## Slide 17: Measurement of Exposure to the X-ray beam N-60 at low dose
Exposure: 0.40 mSv
Lower Limit: 0.18 mSv
Upper Limit: 0.69 mSv

## Slide 18: Thank you for your attention
Office of Atoms for Peace